PSUR Tables — FormQAR-054

Laparoscopic Stapler X100® — Reporting Period Jan-2023 to Dec-2023 (ANNUALLY cadence, EU Class CLASS_IIB)

UDI-DI / Device Identification (Section B)

| Basic UDI-DI | Device Trade Name | EMDN Code | Changes from Previous PSUR |
| --- | --- | --- | --- |
| TD1001 / UDI-DI-LSX100 | Laparoscopic Stapler X100®; Laparoscopic Stapler X100 Reload™ | Not applicable | No changes since previous PSUR (PSUR001). |

Associated Documents (Section B)

| Document Type | Document Number | Document Title |
| --- | --- | --- |
| PMS Plan | PMS-1001R | Post Market Surveillance Report for Laparoscopic Stapler X100 |
| Clinical Evaluation Report | CE1234R, Version 3 | CER for Laparoscopic Stapler X100 (last update 20 October 2023) |
| PMCF Plan | PMCF-1001P | Post-Market Clinical Follow-up Plan for Laparoscopic Stapler X100 |
| Risk Management File | RMF1000-RMR; RMF1000-RA | Risk Management File for Laparoscopic Stapler X100 |
| Technical Documentation | TD1001 | Technical Documentation for Laparoscopic Stapler X100 |
| Instructions for Use | IFU-LSX100 (v5.2) | Instructions for Use — Laparoscopic Stapler X100 |
| Previous PSUR | PSUR001 | Previous Periodic Safety Update Report (ended 2023-03-31) |

Table 1 — Annual Number of Devices Sold by Region

| Region | Preceding 12-Month Period (Jan-2022 to Dec-2022) | Current Data Collection Period (Jan-2023 to Dec-2023) | 12-Month Total | 12-Month Percent of Global Sales |
| --- | --- | --- | --- | --- |
| EEA+TR+XI | Data not available | 5,902 | 5,902 | 49.0% |
| Australia | Data not available | 0 | 0 | 0.0% |
| Brazil | Data not available | 0 | 0 | 0.0% |
| Canada | Data not available | 0 | 0 | 0.0% |
| China | Data not available | 0 | 0 | 0.0% |
| Japan | Data not available | 0 | 0 | 0.0% |
| UK | Data not available | 0 | 0 | 0.0% |
| United States | Data not available | 6,135 | 6,135 | 51.0% |
| Rest of World | Data not available | 0 | 0 | 0.0% |
| Worldwide | Data not available | 12,037 | 12,037 | 100.0% |

Table 2 — Serious Incidents by IMDRF Annex A (MDP) by Region

| Region | IMDRF Problem Code & Term | N (current period) | Rate (%) | Complaint number |
| --- | --- | --- | --- | --- |
| EEA+TR+XI | N/A — No serious incident | 0 | 0.0000% | N/A |
| UK | N/A — No serious incident | 0 | 0.0000% | N/A |
| Worldwide | N/A — No serious incident | 0 | 0.0000% | N/A |

Table 3 — Serious Incidents by IMDRF Annex C (Cause)

| Region | IMDRF Cause Code & Term | N (current period) | Rate (%) | Complaint number |
| --- | --- | --- | --- | --- |
| EEA+TR+XI | N/A — No serious incident | 0 | 0.0000% | N/A |
| UK | N/A — No serious incident | 0 | 0.0000% | N/A |
| Worldwide | N/A — No serious incident | 0 | 0.0000% | N/A |

Table 4 — IMDRF Annex F (Health Impact) × Annex D (Investigation Conclusion)

| IMDRF Health Impact (Annex F) code and term, by region | Number of serious incidents | Conclusion 1 % | Conclusion 2 % | Conclusion 3 % | Conclusion 4 % |
| --- | --- | --- | --- | --- | --- |
| EEA+TR+XI — N/A — No serious incident | 0 | N/A | N/A | N/A | N/A |
| UK — N/A — No serious incident | 0 | N/A | N/A | N/A | N/A |
| Worldwide — N/A — No serious incident | 0 | N/A | N/A | N/A | N/A |

Table 6 — Feedback by Type and Source

| Feedback Type | Source | Count | Summary |
| --- | --- | --- | --- |
| Complaint | End-users (surgeons, hospital staff) | 20 | All complaints summarized in Section F. |
| Non-complaint | Distributors / importers | 0 | No safety-related feedback outside the complaint handling system. |
| Non-complaint | Sales / Customer Service | 0 | No qualitative themes impacting the safety/risk profile of the device. |
| PMCF-derived | PMCF-1001P | 0 | No new signals from PMCF activity (see Section L). |

Table 7 — Complaint Rate by Harm and Medical Device Problem (Annual)

| Harm / Medical Device Problem | Current 12-Month Data Collection Period (Jan-2023 to Dec-2023) | Max Expected Rate of Occurrence (from the RACT) |
| --- | --- | --- |
| Skin/Subcutaneous Injury (Laceration) | 0.0498% (6) | ≤0.01% (O1) |
| Failure to operate as intended (A020101) | 0.0249% (3) | ≤0.01% (O1) |
| Fracture / break of device (A070101) | 0.0166% (2) | ≤0.01% (O1) |
| Packaging defect — incorrect quantity (A0601) | 0.0083% (1) | ≤0.01% (O1) |
| Tissue Reaction (Staple Migration/Extrusion) | 0.0415% (5) | ≤0.01% (O1) |
| Packaging defect — incorrect quantity (A0601) | 0.0166% (2) | ≤0.01% (O1) |
| Failure to operate as intended (A020101) | 0.0166% (2) | ≤0.01% (O1) |
| Fracture / break of device (A070101) | 0.0083% (1) | ≤0.01% (O1) |
| No Health Consequence or Impact | 0.0415% (5) | N/A |
| Fracture / break of device (A070101) | 0.0249% (3) | N/A |
| Packaging defect — incorrect quantity (A0601) | 0.0083% (1) | N/A |
| Foreign material in/on device (A040102) | 0.0083% (1) | N/A |
| Near Miss | 0.0332% (4) | N/A |
| Failure to operate as intended (A020101) | 0.0083% (1) | N/A |
| Fracture / break of device (A070101) | 0.0083% (1) | N/A |
| Packaging defect — incorrect quantity (A0601) | 0.0083% (1) | N/A |
| Foreign material in/on device (A040102) | 0.0083% (1) | N/A |
| Grand Total | 0.1662% (20) | — |

Table 8 — FSCA Initiated in Current Reporting Period

| Type of action | Manufacturer Reference number | Issuing Date / Date of Final FSN | Scope of the FSCA / Device models | Status of the FSCA | Rationale and description | Impacted regions | Date reported to MHRA |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Field Safety Notice | FSCA01 | 2023-06-23 | Laparoscopic Stapler X100 | Completed | Device breakage observed in field (effectiveness: Effective). | NorthAmerica, Europe, Asia | [TO BE COMPLETED] |
| Field Safety Notice | FSCA02 | 2023-08-14 | Laparoscopic Stapler X100 | Completed | Unexpected tissue reaction (effectiveness: Effective). | Europe, NorthAmerica | [TO BE COMPLETED] |
| Field Safety Notice | FSCA03 | 2023-08-15 | Laparoscopic Stapler X100 | Completed | Device breakage observed in field (effectiveness: Pending). | Asia, NorthAmerica, Europe | [TO BE COMPLETED] |

Table 9 — CAPA Initiated in Current Reporting Period

| CAPA Number / Manufacturer Reference | Initiation Date | Scope of the CAPA | Status of the CAPA | CAPA description | Root cause | Effectiveness of the CAPA | Target date for completion |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CAPA01 | 2023 (date not recorded in source) | Laparoscopic Stapler X100 | Closed — effectiveness verified | Trigger: Complaint. Action: Improve packaging. | Mechanical jam | Effective | [TO BE COMPLETED] |
| CAPA02 | 2023 (date not recorded in source) | Laparoscopic Stapler X100 | Closed — effectiveness not demonstrated | Trigger: Trend. Action: Update manual. | Inadequate training | Not Effective | [TO BE COMPLETED] |
| CAPA03 | 2023 (date not recorded in source) | Laparoscopic Stapler X100 | Closed — effectiveness verified | Trigger: Complaint. Action: Retrain staff. | Design flaw | Effective | [TO BE COMPLETED] |
| CAPA04 | 2023 (date not recorded in source) | Laparoscopic Stapler X100 | Open — effectiveness verification underway | Trigger: Audit. Action: Update manual. | Mechanical jam | Pending | [TO BE COMPLETED] |
| CAPA05 | 2023 (date not recorded in source) | Laparoscopic Stapler X100 | Open — effectiveness verification underway | Trigger: Audit. Action: Update manual. | Design flaw | Pending | [TO BE COMPLETED] |

Table 10 — Adverse Events and Recalls (External Databases)

| Database/Registry | Total matches | Relevant findings | Benchmark vs similar devices | Regulatory actions affecting similar devices | RMF update reference |
| --- | --- | --- | --- | --- | --- |
| U.S. FDA MAUDE (product code [TO BE COMPLETED]) | 0 | No adverse events identified for Laparoscopic Stapler X100 during the reporting period. | Competitor devices show similar device-malfunction events (E001, E005, E006, E009). | No FDA recall actions affecting similar competitor devices identified in the period. | RMF1000-RA |
| U.S. FDA Recall Database | 0 | No recalls identified for Laparoscopic Stapler X100 during the reporting period. | No recalls identified affecting comparable stapling devices in the period. | N/A | RMF1000-RA |
| UK MHRA Yellow Card | 0 | No Yellow Card reports identified for Laparoscopic Stapler X100. | No notable signals identified affecting comparable devices. | N/A | RMF1000-RA |
| Australia TGA DAEN | 0 | No DAEN entries identified for Laparoscopic Stapler X100. | No notable signals identified affecting comparable devices. | N/A | RMF1000-RA |
| Health Canada Medical Device Incident Reports | 0 | No incident reports identified for Laparoscopic Stapler X100. | No notable signals identified affecting comparable devices. | N/A | RMF1000-RA |
| EUDAMED | Limited public access | Vigilance module partially available; no signal identified. | 1 entries linked to Laparoscopic Stapler X100 (E007). | No FSCA published for comparable devices in EUDAMED public modules. | RMF1000-RA |