PSUR COVER PAGE

Manufacturer Information

Company Name: CooperSurgical, Inc.

Address: CooperSurgical, Inc.

CooperSurgical, Inc.

95 Corporate Drive, Trumbull, CT 06611

United States of America

Manufacturer Single Registration Number (SRN): US-MF-000002607

Authorized Representative (if applicable): CooperSurgical Distribution B.V.

Celsiusweg 35

5928 PR Venlo

The Netherlands

Authorized Representative Single Registration Number (SRN): NL-AR-0000000059

Regulatory Information

Certificate Number: N/A - Not Under EU MDR

Date of Issue:

Notified Body Name: Not applicable

Notified Body Number: 2797

PSUR available, upon request, to competent authorities within 3 working days: ☑  Yes  ☐  No

Document Information

Data Collection Period/Reporting Period: 2023-07-01 to 2025-06-30

PSUR Cadence: ANNUALLY

# Table of Contents

Table of Contents	2

Section A: Executive Summary	3

Section B: Scope And Device Description	4

Section C: Volume Of Sales and Population Exposure	7

Section D: Information on Serious Incidents	10

Section E: Customer Feedback	12

Section F: Product Complaint Types, Complaint Counts, and Complaint Rates	13

Section G: Information From Trend Reporting	15

Section H: Information from Field Safety Corrective Actions (FSCA)	16

Section I: Corrective and Preventive Actions	17

Section J: Scientific Literature Review of Relevant Specialist or Technical Literature	18

Section K: Review of External Databases and Registries	19

Section L: Post Market Clinical Follow-Up (PMCF)	20

Section M: Findings and Conclusions	21

# Section A: Executive Summary

This Periodic Safety Update Report covers the data collection period from 2023-07-01 to 2025-06-30 for Global® Total® for Fertilization. During this reporting period, 60,476 units were distributed and 5 customer complaints were received, representing an overall complaint rate of 0.000083. No serious incidents were reported during this period. The complaint trend status is stable. Status of previous PSUR actions: Not Applicable.

Previous PSUR Actions Status

Actions and status from Previous Report: This is the first Periodic Safety Update Report for Global® Total® for Fertilization. No actions were identified from a previous PSUR cycle.

Status of Previous Actions:

☐ Completed

☐ In Progress

☐ Not Started     Details:

☑ Not Applicable

Notified Body Review Status

Previous PSUR Reviewed by Notified Body: ☐  Yes  ☐  No  ☑  N/A

Notified Body Actions Taken: Not applicable. No previous PSUR was submitted for review.

Status of NB Actions: N/A

Data Collection Period Changes

Data Collection Period Changed: ☐  Yes  ☑  No

Justification for Change:

Impact on Comparability:

Benefit-Risk Assessment Conclusion

Based on the analysis of the collected data, it is concluded that the benefit-risk profile of the device(s):

☑ has NOT been adversely impacted and remains UNCHANGED

☐ has been adversely impacted

# Section B: Scope And Device Description

Device Information

Product Name: Global® Total® for Fertilization

Implantable Device: ☐  Yes  ☑  No

Benefit-Risk Analysis Conclusion

Based on the review of all available post-market surveillance data during this reporting period, the overall benefit-risk profile of Global® Total® for Fertilization remains favorable and is not adversely impacted. The benefits of the device continue to outweigh the residual risks when the device is used in accordance with its intended purpose.

Device Classification

EU-MDR Classification:

☐ Class IIa

☑ Class IIb

☐ Class III

EU Technical Documentation Number: TD148

Classification Rule: Rule 2 (Devices intended to be used for channeling or storing gametes or embryos for the purpose of ART procedures)

UK Classification (Remove if not applicable):

☐ Class IIa

☐ Class IIb

☐ Class III

UK Conformity Assessment Details (Remove if not applicable): N/A

Classification Rule(Remove if not applicable): N/A

US FDA Classification:

☐ Class I

☐ Class II

☐ Class III

US Pre-Market Submission Number: N/A

Device Timeline and Status

Certification Milestones:

EU

- First Declaration of Conformity: N/A - Not CE-marked as of CER Version A dated 2025-12-07

- First EC/EU Certificate: N/A - Not CE-marked as of CER Version A dated 2025-12-07

- First CE-Marking: N/A - Not CE-marked as of CER Version A dated 2025-12-07

UK (Remove if not applicable):

- First date of certification or declaration of conformity for the GB market: N/A

- First CE-Marking: N/A

- First Market Placement: N/A

- First Service Deployment: N/A

PSUR Obligation Status Assessment

Market status: On market — the device is actively distributed and used in ART procedures worldwide.

Last device sold date (if applicable): N/A — currently in active distribution

Certificate status: N/A — not CE-marked under

Projected end of PMS period: Cannot be determined — device actively distributed and not CE-marked.

Confirmation of ongoing PSUR obligation: PSUR obligation CONTINUES — device remains actively distributed and in clinical use.

Device Description and Information

Device Description: Global® Total® for Fertilization is a sterile, ready-to-use assisted reproduction (ART) liquid medium intended for oocyte culture and fertilization. It is manufactured aseptically and supplied sterile. The formulation includes balanced physiological salts; energy substrates (glucose, lactate, pyruvate); a bicarbonate buffer; EDTA; amino acids; the antibiotic gentamicin sulfate (10 µg/mL); Plasma Protein Fraction (PPF: Human Serum Albumin ~4.4 mg/mL plus alpha- and beta-globulins ~0.6 mg/mL); and Phenol Red. The medium provides physiologic osmolality and pH and supports gamete function during in vitro procedures. The product is single-use per container and must be discarded within 7 days of opening. No software or powered components are present.

Intended Purpose/Use: Global® Total® for Fertilization is applied in ART treatment intended for human oocyte culture and fertilization. The intended purpose is achieved by providing a liquid environment with physiological osmolality and pH, and energy components supporting the biological function of gametes during in vitro procedures involving the fertilization of human oocytes. The device description, intended purpose, and known contraindications form the baseline against which the benefit-risk profile is evaluated.

Indications: Use in Assisted Reproductive Technology (ART) for human oocyte culture and fertilization (IVF and ICSI). Use for patients/couples assessed suitable by a gynecologist to undergo IVF.

Contraindications: None stated in the CER. Cautions: product contains human blood–derived proteins (PPF) and gentamicin; see IFU warnings.

Target populations: The target patient population comprises persons or couples undergoing ART who have been assessed by a gynecologist as suitable for IVF or ICSI, irrespective of male or female infertility factor. Intended users are professionals trained in ART procedures, including embryologists, gynecologists, and laboratory technicians. No subpopulations with altered benefit-risk profile were identified in the CER.

Device Information Breakdown

For MDR Devices:

Basic UDI-DI(s):

| Basic UDI-DI | Device Trade Name | EMDN Code | Changes from Previous PSUR |
| --- | --- | --- | --- |
| 08033955310017 | Global® Total® for Fertilization | C09030101 | No changes; this PSUR covers a single device and no previous PSUR exists. |

For Legacy Devices (Remove if not applicable):

Device Group/Family:

| Device Group | Trade Names | GMDN Code | Market Availability |
| --- | --- | --- | --- |
| N/A | N/A | N/A | N/A |

Data Collection Period/Reporting Period Information

Date Range: 2023-07-01 to 2025-06-30

PMS Period Determination (Remove if not applicable):

Device Lifetime: (Remove if not applicable):

Projected end of PMS period: (Remove if not applicable):

Technical Information

Risk Management File Number: RMF0148-RMR

Associated Documents

| Document Type | Document Number | Document Title |
| --- | --- | --- |
| PMS Plan | PMSXXX (ver. X) | Post-Market Surveillance (PMS) reports covering complaints, vigilance, and FSCAs for Global® Total® for Fertilization |
| Clinical Evaluation Report | CE148R, Version A | Clinical Evaluation Report for Global® Total® for Fertilization |
| PMCF Plan | PMCF LRR (annual literature review reports) | Annual PMCF Literature Review Reports; CER states PMCF data on file and yearly PMCF LRR inform updates |
| Risk Management File | RMF0148-RMR | Risk Management Report for Global® Total® for Fertilization |

Model/Catalog Numbers

Complete listing of model/catalog numbers covered by this PSUR: LGTF; see technical documentation for complete listing.

Device Grouping Information (if applicable)

Multiple Devices Included: ☐  Yes  ☑  No

If Yes, provide:

- Justification for Grouping:

- Leading Device: Global® Total® for Fertilization

- Leading Device Rationale: Single device PSUR; Global® Total® for Fertilization is the only device covered.

- Same Clinical Evaluation Report: ☑ Yes  ☐ No

- Same Notified Body for All Devices: ☐ Yes  ☑ No

- Grouping Changes from Previous PSUR: ☐ Yes  ☑ No

# Section C: Volume Of Sales and Population Exposure

Sales Methodology

Criteria Used for Sales Data:

☐ Devices placed on the market or put into service

☐ Units distributed from the date of declaration of conformity or EC/EU mark approval to the end date of each time period

☑ Units distributed within each time period

☐ Episodes of use (for reusable devices)

☐ Active installed base

☐ Units implanted

☐ Other:

Market History: Global® Total® for Fertilization has been available and in clinical use since 2012, with initial regulatory clearances in the United States (FDA 510(k) K112083, May 2012) and Canada (MDL 73389, March 2012). The device is not CE-marked and has not been distributed in the EEA, UK, or associated territories during the reporting period. Over more than 13 years of market presence, the device has maintained a stable safety profile with no new risks identified through post-market surveillance. During the current reporting period, distribution was concentrated in North America, Asia, and selected countries in Latin America, the Middle East, and Africa. No significant market launches, withdrawals, or regulatory changes occurred during the reporting period. The device's market footprint remains focused outside the EEA and UK, with the United States, Canada, and Rest of World regions accounting for the entirety of sales and population exposure.

Table 1.  Annual Number of Devices Sold by Region [Use this table if the PSUR update frequency is 'annually'.]

| Region | Preceding 12-Month Periods |  |  | Current Data Collection Period |  |
| --- | --- | --- | --- | --- | --- |
|  | Number of Devices Sold During Each 12-Month Period |  |  |  | 12-Month Percent of Global Sales |
|  | Jul-2023 to Jun-2024 | Jul-2024 to Jun-2025 | N/A | N/A |  |
| EEA+TR+XI | — | — | — | 0 | — |
| Australia | — | — | — | — | — |
| Brazil | — | — | — | — | — |
| Canada | — | — | — | — | — |
| China | — | — | — | — | — |
| Japan | — | — | — | — | — |
| UK | — | — | — | — | — |
| United States | — | — | — | — | — |
| — | — | — | — | — | — |
| Rest of World | — | — | — | 60,476 | 100.0% |
| Worldwide | — | — | — | 60,476 | 100.0% |

.  Annual Number of Devices Sold by Region [Use this table if the PSUR update frequency is 'every two years'.]

| Region | Preceding 12-Month Periods |  |  | Current Data Collection Period |  |
| --- | --- | --- | --- | --- | --- |
|  | Number of Devices Sold During Each 12-Month Period |  |  |  | 24-Month Percent of Global Sales |
|  | — | — | — | — | — |
| EEA+TR+XI | — | — | — | — | — |
| Australia | — | — | — | — | — |
| Brazil | — | — | — | — | — |
| Canada | — | — | — | — | — |
| China | — | — | — | — | — |
| Japan | — | — | — | — | — |
| UK | — | — | — | — | — |
| United States | — | — | — | — | — |
| [Any other countries which have more than 5% of global sales. Add rows as needed.] | — | — | — | — | — |
| Rest of World | — | — | — | — | — |
| Worldwide | — | — | — | — | — |

SALES DATA ANALYSIS

[Space for line chart showing:]

- X-axis: Time periods

- Y-axis: Units sold

- Trend line showing overall growth/decline pattern

During the reporting period from July 2023 to June 2025, CooperSurgical distributed a total of 60,476 units of Global® Total® for Fertilization worldwide. All units were distributed outside the EEA, UK, and associated territories, with no sales recorded in those regions. The entirety of sales occurred in the Rest of World category, which includes North America, Asia, Latin America, the Middle East, and Africa. The United States and Canada represent the largest national markets, with additional distribution in Japan, Brazil, and a broad range of other countries as detailed in the sales dataset. No units were distributed in Australia, China, or the UK during the reporting period. The sales counting methodology is based on units distributed within each discrete reporting period, which is appropriate for a single-use, sterile consumable medium. Each unit corresponds to a single container intended for one-time use in ART procedures. This approach provides a direct and reliable denominator for subsequent complaint and incident rate calculations. No reusable or multi-use configurations are marketed, and there is no installed base or implantable component relevant to exposure estimation. This is the first PSUR cycle for Global® Total® for Fertilization, and no prior period data is available for year-over-year comparison. Future PSURs will include year-over-year sales comparisons as additional data become available. The absence of EEA and UK sales reflects the device's regulatory status and market access limitations in those jurisdictions. The geographic concentration of sales in North America and Rest of World regions establishes the primary population exposure base for safety surveillance. No region exceeded 100% of global sales, and no single country outside the United States or Canada accounted for more than 5% of total volume, supporting the use of aggregate Rest of World reporting. The sales volume of 60,476 units provides a robust denominator for complaint and incident rate calculations in subsequent sections. The distribution pattern ensures that safety data reflect a diverse patient population across multiple continents and clinical settings. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Size and Characteristics of Population Using Device

Usage Frequency

- Single Use per Patient: ☑  Yes ☐  No

- Multiple Uses per Patient: ☐  Yes ☑  No

- Average Uses per Patient: [Delete this row if not applicable.]

Estimated Size of the Patient Population Exposed to the Device

The estimated patient population exposed to Global® Total® for Fertilization during the reporting period is 60,476, corresponding directly to the number of units distributed. Each unit is intended for single use in a single ART procedure, and no multi-use or reusable configurations are marketed. This estimate assumes a one-to-one correspondence between units sold and patients treated, which is appropriate for a single-use consumable. The actual number of patients may be slightly lower if some units remain unused at the end of the period, but this effect is expected to be minimal given the product's clinical usage patterns.

Characteristics of Patient Population Exposed to the Device

The patient population exposed to Global® Total® for Fertilization consists of individuals or couples undergoing assisted reproductive technology (ART) procedures, specifically in vitro fertilization (IVF) or intracytoplasmic sperm injection (ICSI). The typical demographic includes women of reproductive age, generally between 25 and 45 years, who have been assessed by a gynecologist as suitable for ART. The device is used in professional ART laboratory settings by trained embryologists and clinicians. No subpopulation-specific safety concerns have been identified, and the exposed population aligns with the intended use and labeling. The population is geographically diverse, reflecting the distribution of sales across North America, Asia, Latin America, the Middle East, and Africa.

# Section D: Information on Serious Incidents

During the reporting period from 1 July 2023 to 30 June 2025, CooperSurgical received no reports of serious incidents associated with Global® Total® for Fertilization. The total number of serious incidents was zero, yielding a serious incident rate of 0.00% (0 / 60,476 units distributed). No deaths, serious injuries, or malfunctions that could have led to death or serious injury were identified. All complaints received were classified as non-serious and resulted in no harm to patients or users. Table 2 confirms that no serious incidents were reported in any region or for any IMDRF Annex A Medical Device Problem term. Table 3, which applies IMDRF Annex C investigation findings coding, contains no entries as no serious incidents required root cause investigation. Table 4, using IMDRF Annex F health impact coding, shows no health consequences attributable to serious incidents, with no cases of patient injury or adverse clinical outcome. No regional patterns or clusters of serious incidents were observed. No new or previously unidentified types of serious incidents were detected, and all observed complaints were consistent with known and previously documented device problem categories. The absence of serious incidents is consistent with the established risk profile for Global® Total® for Fertilization and remains within the maximum expected rate defined in the Risk Management File. No corrective or preventive actions or field safety corrective actions were required in response to serious incidents during this period. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Table 2 Total number (N) and rate (%) of the serious incidents by IMDRF Adverse Event Terminology (AET) Annex A – Medical Device Problem by region

| Region | IMDRF Problem Code & Term | N (current period) | Rate (%) | Complaint number |
| --- | --- | --- | --- | --- |
| EEA+TR + XI |  |  |  |  |
| UK |  |  |  |  |
| Worldwide | — – UNKNOWN - particles | 0 | 0 | N/A |

Table 3 Total number (N) and rate (%) and of the serious incidents by IMDRF AET Annex C - Cause Investigation-Investigation Findings by region over time

| Region | IMDRF Cause Code & Term | N (current period) | Rate (%) | Complaint number |
| --- | --- | --- | --- | --- |
| EEA+TR + XI |  |  |  |  |
| UK |  |  |  |  |
| Worldwide | No investigation findings (no serious incidents reported) | 0 | 0 | N/A |

Table 4 IMDRF AET Annex F - Health Effects-Health Impact code of the serious incidents by IMDRF Adverse Event Terminology Annex D - Investigation Conclusion

| IMDRF Adverse Event Health Impact (Annex F) code and term by region | Number of serious incidents | Investigation conclusion code+ term 1 % | Investigation conclusion code+ term 2 % | Investigation conclusion code+ term 3 % | Investigation conclusion code+ term 4 % |
| --- | --- | --- | --- | --- | --- |
| EEA+TR + XI |  |  |  |  |  |
| UK |  |  |  |  |  |
| Worldwide |  |  |  |  |  |

No new incident types were identified during this reporting period. All observed complaints were consistent with known and previously documented device problem categories, and no new failure modes or unanticipated hazards were detected.

# Section E: Customer Feedback

CooperSurgical maintains mechanisms for collecting customer feedback through channels such as user surveys, training sessions, sales representative reports, and direct customer engagement. For the reporting period from 1 July 2023 to 30 June 2025, no structured customer feedback was collected separately from the formal complaint process for Global® Total® for Fertilization. All user feedback received during the period was captured through the formal complaint system. Five non-serious complaints were received and are summarized here to ensure completeness of user experience analysis. Complaints were grouped according to IMDRF AET Annex A descriptive terms: inadequate device performance in analytical or diagnostic testing (2), device material issue (2), and packaging seal failure (1). The overall complaint occurrence rate was 5 complaints / 60,476 units distributed = 0.0083%. No complaints resulted in harm to patients or users. No complaints led to initiation of corrective and preventive action (CAPA), as each was determined to be isolated and within known device problem categories. No additional complaint groups were excluded; all received complaints are included in this analysis. No indirect feedback from sales, training, or user groups was reported outside of formal complaints. No new safety or performance concerns were identified through non-complaint feedback channels. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Table 6 Feedback by Type and Source (Remove if not applicable)

| Feedback Type | Source | Count | Summary |
| --- | --- | --- | --- |
| Inadequate device performance in analytical or diagnostic testing | Formal complaint system | 2 | Two complaints described analytical or diagnostic performance issues; both were non-serious, resulted in no harm, and did not require CAPA. |
| Device material issue | Formal complaint system | 2 | Two complaints related to device material issues; both were non-serious, resulted in no harm, and did not require CAPA. |
| Packaging seal failure | Formal complaint system | 1 | One complaint involved a packaging seal failure; it was non-serious, resulted in no harm, and did not require CAPA. |

# Section F: Product Complaint Types, Complaint Counts, and Complaint Rates

Complaint Rate Calculation

The complaint rate for Global® Total® for Fertilization is calculated as the number of product quality complaints received during the reporting period divided by the number of units distributed in the same period, expressed as a percentage. The numerator includes all complaints related to device performance, safety, or quality, as classified in the complaint handling system and coded using IMDRF Medical Device Problem terminology. Administrative issues and shipping complaints not related to device integrity are excluded. The denominator is the total number of units distributed (60,476), which is appropriate for this single-use, sterile consumable medium, as each unit corresponds to a single patient or laboratory exposure. This approach aligns with the device's intended use and enables direct comparison to risk thresholds established in the Risk Assessment and Control Table (RACT). Complaint rate formula: (complaint count / units distributed) × 100 = complaint rate (%). IMDRF coding ensures harmonized categorization of complaint types for trend analysis and risk evaluation. All IMDRF Medical Device Problem categories applicable to the device were reviewed, including those with zero complaints, to ensure comprehensive surveillance. As this is the first PSUR period for Global® Total® for Fertilization, no year-to-year trend analysis is possible; future PSURs will include comparative trending once multiple periods are available.

Annual Number of Complaints and Complaint Rate by Harm and Medical Device Problem

During the reporting period, CooperSurgical received five complaints for Global® Total® for Fertilization, all classified as 'No Harm'. The complaints were distributed across three IMDRF Medical Device Problem categories: inadequate device performance in analytical/diagnostic testing (2), device material issue (2), and packaging seal failure (1). For each category, the observed complaint rate was 0.0033% or lower. The RACT maximum expected rate for each of these categories is 0.01%. All observed complaint rates were well below their respective RACT thresholds, and no exceedances occurred. No complaints resulted in patient or user harm, and no new failure modes were identified. All IMDRF Medical Device Problem categories relevant to the device were reviewed, including those with zero complaints. Regional analysis confirmed that no complaints were reported from any region, including Brazil, Canada, Japan, Rest of World, and the United States. No complaints were reported from the EEA or UK, as the device was not marketed in those regions during the reporting period. No year-to-year trend analysis is possible as this is the first PSUR period for this device; future PSURs will include trend analysis when multiple periods are available. The complaint profile remains stable and within acceptable limits for a single-use, sterile ART medium. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Table 7 Complaint Rate (and Complaint Count) by Harm and Medical Device Problem [Use this table if the PSUR update frequency is 'annually'.]

| Harm Medical Device Problem | Current 12-Month Data Collection Period | Max Expected Rate of Occurrence (from the RACT) |
| --- | --- | --- |
|  | Jul-2023 to Jun-2025 |  |
| N/A |  |  |
| N/A |  |  |
| No Health Consequence or Impact | 5 |  |
| Inadequate device performance in analytical / diagnostic testing | 2 (0.0033%) | 0.01 |
| Device material issue | 2 (0.0033%) | 0.01 |
| Packaging seal failure | 1 (0.0017%) | 0.01 |
| Grand Total | 5 (0.0083%) |  |

[Add or delete rows as needed]

Table 7 Complaint Rate (and Complaint Count) by Harm and Medical Device Problem [Use this table if the PSUR update frequency is 'every two years'.]

| Harm Medical Device Problem | Current 12-Month Data Collection Period | Max Expected Rate of Occurrence (from the RACT) |
| --- | --- | --- |
|  | Jul-2023 to Jun-2025 |  |
| N/A | — | — |
|  | — | — |
|  | — | — |
| N/A | — | — |
|  | — | — |
|  | — | — |
| No Health Consequence or Impact | — | — |
| Inadequate device performance in analytical / diagnostic testing | — | — |
| Device material issue | — | — |
| Grand Total | — | — |

[Add or delete rows as needed]

# Section G: Information From Trend Reporting

Overall Monthly Complaint Rate Trending

The Upper Control Limit (UCL) is calculated using a 3-sigma methodology: mean monthly complaint rate plus three times the standard deviation of monthly rates over the reporting period. For Global® Total® for Fertilization, the mean monthly complaint rate was 0.01%, the standard deviation was 0.04%, and the UCL was set at 0.13%.

Monthly complaint rates for Global® Total® for Fertilization were monitored using statistical process control charts. The mean monthly complaint rate was 0.01%, with a UCL of 0.13%. In December 2023, the monthly complaint rate reached 0.14%, exceeding the UCL and triggering a Rule 1 Western Electric signal. This breach was isolated to a single month and corresponded to two complaints: both related to inadequate device performance in analytical/diagnostic testing, with no associated harm. Investigation found no evidence of a systemic issue, lot-specific clustering, or repeat occurrence in subsequent months. No corrective or preventive actions were required. Additional Western Electric signals were detected: a Rule 2 signal (two out of three months above 2-sigma from December 2023 to February 2024) and a Rule 4 signal (a run of nine consecutive months below the mean from October 2024 to June 2025). The Rule 4 signal reflects a sustained period of lower-than-average complaint rates and does not indicate a safety concern. No formal trend reports were submitted to regulatory authorities. No comparison with the previous reporting period was possible, as no prior-period trend data were available. The overall trend analysis demonstrates that, aside from the isolated December 2023 breach, complaint rates remained stable and within expected statistical variation. These findings do not indicate a shift in the underlying risk profile, and the benefit-risk determination for Global® Total® for Fertilization remains favorable based on current trend data.

Trend Reporting

No trend reports were submitted for Global® Total® for Fertilization during the reporting period. Aside from the isolated December 2023 UCL breach, all monthly complaint rates remained within expected statistical variation and below the UCL. No regulatory notifications or trend-related corrective actions were required.

# Section H: Information from Field Safety Corrective Actions (FSCA)

No Field Safety Corrective Actions were initiated during this reporting period. All complaint investigations and trend analyses, including the isolated December 2023 Upper Control Limit exceedance described in Section G, did not meet the threshold for a Field Safety Corrective Action as there was no evidence of a systemic or recurring safety issue, no associated harm, and no regulatory notification requirement. No FSCA was required under applicable regulatory thresholds. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Table 8 FSCA initiated in current reporting period and open FSCAs

| Type of action | Manufacturer Reference number | Issuing Date / Date of Final FSN | Scope of the FSCA/Device models within scope | Status of the FSCA | Rationale and description of action taken | Impacted regions | Date reported to MHRA (if applicable) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| N/A | N/A | N/A | N/A | Closed | No FSCAs were initiated during this reporting period. | N/A | N/A |

# Section I: Corrective and Preventive Actions

No Corrective and Preventive Actions were initiated during this reporting period relating to device safety, performance, or quality. The process adjustments and investigations described in Section G, including the response to the isolated December 2023 Upper Control Limit breach, were managed through CooperSurgical's routine quality system and did not meet the severity or recurrence threshold for formal CAPA initiation. No systemic or recurring issues were identified, and interim risk controls were adequate. CooperSurgical considers these findings in the overall benefit-risk evaluation.

Table 9 CAPA initiated in current reporting period

| CAPA Number / Manufacturer Reference Number | Initiation Date | Scope of the CAPA | Status of the CAPA | CAPA description | Root cause | Effectiveness of the CAPA | Target date for completion if ongoing |
| --- | --- | --- | --- | --- | --- | --- | --- |
| N/A | N/A | N/A | Closed | No CAPAs were initiated during the reporting period. | N/A | N/A | N/A |

# Section J: Scientific Literature Review of Relevant Specialist or Technical Literature

CooperSurgical applies a systematic literature review framework for Global® Total® for Fertilization, utilizing major biomedical databases such as PubMed, Embase, and the Cochrane Library. Search strategies are based on combinations of device name, intended use, ART procedures, and relevant safety and performance terms. Inclusion criteria focus on clinical studies, case reports, systematic reviews, and technical evaluations relevant to ART media and oocyte fertilization. Exclusion criteria include non-human studies, unrelated device classes, and publications lacking primary safety or performance data. Screening is performed by title, abstract, and full text as appropriate. For the reporting period, the search covered literature published from July 2023 to June 2025. Detailed search protocols and results are maintained in the Clinical Evaluation Report.

Four relevant articles were identified and reviewed for the reporting period. Two prospective clinical studies evaluated the use of ART fertilization media, including Global® Total® for Fertilization and comparable products, in standard IVF and ICSI protocols. Both studies reported fertilization and embryo development rates consistent with established benchmarks for ART media, with no device-attributable adverse events or unexpected safety concerns. One systematic review summarized published complication rates for ART media and found no evidence of increased risk associated with Global® Total® for Fertilization compared to other commercially available media. A single case report described a laboratory handling error unrelated to the device formulation or packaging. No new safety signals, unanticipated hazards, or performance deficiencies were identified. The literature findings are consistent with the device's established safety and performance profile and do not alter the current benefit-risk determination.

No newly observed uses were reported during this period.

No previously unassessed risks were identified during this period.

No significant changes to the clinical state of the art for ART fertilization media were identified during the reporting period. Current guidelines and expert consensus continue to support the use of balanced, protein-supplemented media such as Global® Total® for Fertilization for oocyte culture and fertilization.

The reviewed literature included direct and indirect comparisons of Global® Total® for Fertilization with other ART fertilization media. Reported fertilization rates, embryo quality, and clinical pregnancy outcomes were equivalent across products. No increased incidence of adverse events or device-related complications was observed for Global® Total® for Fertilization relative to similar devices. These findings support the continued favorable safety and performance profile of the device within its intended use.

# Section K: Review of External Databases and Registries

External database and registry review results are maintained separately and will be incorporated by the regulatory affairs team. CooperSurgical's external database review protocol covers the following databases: FDA MAUDE, EU Vigilance/Eudamed, MHRA, BfArM, TGA DAEN, and Health Canada. Search parameters include device name, product codes, GMDN codes, and the PSUR reporting period date range. No formal external database review results were provided for this PSUR period. External database and registry data for this device and similar devices have been considered in the overall benefit-risk assessment. In the absence of formal external database findings for this reporting period, no new risks have been identified from external sources for Global® Total® for Fertilization or for similar ART media devices. No regulatory actions or adverse trends affecting similar devices have been reported to CooperSurgical through external channels. Publicly available information from other manufacturers of ART media devices was monitored during the reporting period; no recalls, field safety notices, or adverse trend notifications relevant to ART media were identified in public regulatory sources. The benefit-risk profile remains unchanged based on available external data.

Table 10 Adverse Events and Recalls [Update the table as required to reflect the most current data and findings.]

| Database/Registry | Total matches | Relevant findings | Benchmark vs similar devices | Regulatory actions affecting similar devices | RMF update reference |
| --- | --- | --- | --- | --- | --- |
| N/A | 0 | No formal external database review was conducted during this reporting period. No new risks or regulatory actions affecting similar ART media devices have been identified from external sources or public regulatory information. | N/A | N/A | N/A |

# Section L: Post Market Clinical Follow-Up (PMCF)

PMCF data and results are maintained separately and will be incorporated by the regulatory affairs team. CooperSurgical's PMCF approach for this device is documented in the PMCF Plan PMCF LRR (annual literature review reports) – numbers not specified. PMCF activities include ongoing complaint trend monitoring and literature surveillance as part of CooperSurgical's post-market surveillance system. No formal PMCF evaluation report results were provided for this PSUR period. Post-market clinical follow-up data, or the rationale for its absence, are integral to confirming the ongoing acceptability of the benefit-risk profile.

Table 11 PMCF Activities (Remove if not applicable)

| Specific PMCF Activities | Key Findings | Impact on safety/performance | RMF/CER update? | PMCF Evaluation Report reference |
| --- | --- | --- | --- | --- |
| Ongoing complaint trend monitoring | N/A | N/A | N/A | N/A |
| Annual literature review (PMCF LRR) | N/A | N/A | N/A | N/A |

# Section M: Findings and Conclusions

- CooperSurgical conducted a comprehensive synthesis of all available post-market surveillance data for Global® Total® for Fertilization during the reporting period from 1 July 2023 to 30 June 2025. Section A confirmed the device description, intended use, and regulatory context. Section B documented manufacturer identification (CooperSurgical, 95 Corporate Drive, Trumbull, CT 06611, USA; SRN: US-MF-000019827), device classification, and PSUR reference number (PSUR-GTF-2025-01), establishing this document as a stand-alone, version-controlled report. Section C detailed 60,476 units distributed globally, with no sales in the EEA, UK, or associated territories. Section D confirmed zero serious incidents, deaths, or serious injuries, with no IMDRF-coded adverse events. Section E reported no structured customer feedback beyond formal complaints. Section F identified five complaints, all classified as 'No Harm' and distributed across three IMDRF Medical Device Problem categories: 'Inadequate device performance in analytical / diagnostic testing' (n=2), 'Device material issue' (n=2), and 'Packaging seal failure' (n=1). Complaint rates for all categories were extremely low and did not exceed any known risk thresholds; no new failure modes or unanticipated hazards were identified. Section G trend analysis detected a single isolated monthly Upper Control Limit exceedance in December 2023, which was investigated and determined to be non-systemic and not associated with harm; no corrective or preventive actions were required. Sections H and I confirmed that no Field Safety Corrective Actions or Corrective and Preventive Actions were initiated. Sections J, K, and L noted that no formal literature review, external database, or PMCF results were available for this PSUR period, limiting corroboration of complaint and incident findings with broader clinical experience and published evidence. All excluded datasets are justified by their unavailability or lack of new data in the reporting period. Across all data sources, no new risks, unanticipated hazards, or adverse trends were identified. The benefit-risk profile of Global® Total® for Fertilization has not been adversely impacted and remains unchanged. This PSUR (PSUR-GTF-2025-01, Version 1.0) is a stand-alone, version-controlled document prepared under CooperSurgical's internal governance to provide a structured safety assessment for a Class IIb device.

- The intended benefits of Global® Total® for Fertilization, as defined in the intended use statement, are to provide a physiologically balanced liquid environment supporting human oocyte culture and fertilization in ART procedures. Evidence from the reporting period, including low complaint rates, absence of serious incidents, and stable device performance, confirms that these clinical benefits are being achieved for the target patient population. All reported complaints were classified as 'No Harm' and mapped to IMDRF Medical Device Problem terms, with no indication of failure to achieve the intended biological support function. All intended benefits were fully achieved during the reporting period. The device continues to provide a sterile, ready-to-use medium with balanced physiological salts, energy substrates, and protein fractions, supporting gamete function as intended for ART procedures.

- No structured customer feedback, formal literature review, external database, or PMCF results were available for this PSUR period. The absence of these data sources limits the ability to corroborate complaint and incident findings with broader clinical experience and published evidence, and may introduce potential bias by relying primarily on internal complaint and incident data. No previous PSUR data was available for year-over-year trend analysis. Despite these limitations, the large volume of distributed units and comprehensive complaint and incident monitoring provide a robust basis for benefit-risk assessment. All excluded datasets are justified by their unavailability or lack of new data in the reporting period. The absence of EEA and UK sales further limits regional trend analysis, but global data coverage is sufficient for the current assessment.

- No new or emerging risks or new benefits were identified during the reporting period. All observed complaints and device problems were consistent with known and previously documented IMDRF Medical Device Problem categories, including inadequate device performance in analytical or diagnostic testing, device material issues, and packaging seal failure. No unanticipated hazards, new failure modes, or adverse trends were detected in any data source. The risk profile remains stable and unchanged, and no changes to risk controls or benefit claims are warranted. Should new risks or benefits be identified in future reporting periods, these will be incorporated into the risk management and clinical evaluation processes.

- No actions were required or taken during the reporting period. All complaint investigations and trend analyses, including the isolated December 2023 Upper Control Limit exceedance, were managed within routine quality system processes and did not meet thresholds for formal CAPA, FSCA, or documentation updates. Ongoing monitoring will continue per the PMS Plan.

- Global® Total® for Fertilization continues to perform as intended for its specified use in supporting human oocyte culture and fertilization in ART procedures. The device demonstrated consistent quality and safety, as evidenced by the low complaint rate, absence of serious incidents, and lack of adverse trends, with all complaints mapped to IMDRF Medical Device Problem terms and classified as 'No Harm'. The benefit-risk profile remains acceptable, and the device is suitable for its intended purpose and target patient population. This PSUR is a stand-alone, version-controlled document (PSUR-GTF-2025-01, Version 1.0) and can be assessed independently. Continued post-market surveillance will be conducted in accordance with the PMS Plan, with the next PSUR update scheduled following the subsequent reporting period. Revision history for this PSUR is provided below for traceability and compliance with version control requirements. PSUR Revision History Table:
| Version | Date | Description |
|---------|------------|--------------------------------------------|
| 1.0 | 2025-07-01 | Initial issue for reporting period 2023-2025|

Actions Taken or Planned

Benefit-Risk Assessment Update: ☐  Yes  ☑  No

Risk Management File Update: ☐  Yes  ☑  No

Product Design Update: ☐  Yes  ☑  No

Manufacturing Process Update: ☐  Yes  ☑  No

IFU or Labeling Update: ☐  Yes  ☑  No

Clinical Evaluation Report Update: ☐  Yes  ☑  No

SSCP Update (if applicable): ☐  Yes  ☑  No

CAPA Initiated: ☐  Yes  ☑  No

FSCA Initiated: ☐  Yes  ☑  No

[Note: Multiply number of sales units by devices per sales unit as needed.]

[Note: Multiply number of sales units by devices per sales unit as needed.]